[image: image1.jpg]GOOD WILL INSTRUMENT CO.,, LTD.

NO. 7-1, Jhongsing Road, Tucheng Dist., New Taipei City, 236, Taiwan

GE lnerK T (886) 22268-0389 F (886)2 2268-0639  www.gwinstek.com






GSP-9330 Specifications

The specifications apply when the GSP-9330 is powered on for at least 30 minutes under +20°C~+30°C.
	Frequency

	Frequency
	Range
	9 kHz to 3.25 GHz

	
	Resolution
	1 Hz

	Frequency Reference
	Accuracy
	±(period since last adjustment X aging rate) + stability over temperature + supply voltage stability

	
	Aging Rate
	±1 ppm max.
	1 year after last adjustment

	
	Frequency Stability over Temperature
	±0.025 ppm
	0 to 50 °C

	
	Supply Voltage Stability
	±0.02 ppm
	

	Frequency Readout Accuracy
	Start, Stop, Center, Marker
	±(marker frequency indication X frequency reference accuracy + 10% x RBW + frequency resolution)

	
	Trace points
	Max 601 points, min 6 points

	Marker Frequency Counter
	Resolution
	1 Hz, 10 Hz, 100 Hz, 1 kHz

	
	Accuracy
	±(marker frequency indication X frequency reference accuracy + counter resolution)
	RBW/Span >=0.02 ;

Mkr level to DNL>30 dB

	Frequency Span
	Range
	0 Hz (zero span), 100 Hz to 3.25 GHz
	

	
	Resolution
	1 Hz
	

	
	Accuracy
	± frequency resolution1
	RBW: Auto;

	Phase Noise
	Offset from Carrier
	
	Fc =1 GHz; RBW = 1 kHz, VBW = 10 Hz; Average ≥ 40

	
	10 kHz
	<-88 dBc/Hz
	Typical

	
	100 kHz
	<-95 dBc/Hz
	Typical

	
	1 MHz
	<-113 dBc/Hz
	Typical

	Resolution Bandwidth (RBW) Filter
	Filter Bandwidth
	1 Hz to 1 MHz in 1-3-10 sequence
	-3dB bandwidth

	
	
	200 Hz, 9 kHz, 120 kHz, 1MHz
	-6dB bandwidth

	
	Accuracy
	± 8%, RBW = 1MHz
	Nominal

	
	
	± 5%, RBW < 1MHz
	Nominal

	
	Shape Factor
	< 4.5:1
	Normal Bandwidth ratio: -60dB:-3dB

	Video Bandwidth (VBW) Filter
	Filter Bandwidth
	1 Hz to 1 MHz in 1-3-10 sequence
	-3dB bandwidth

	Amplitude

	Amplitude Range
	Measurement Range
	100 kHz to 1 MHz
	Displayed Average Noise Level (DANL) to 18 dBm

	
	
	1 MHz to 10 MHz
	DANL to 21 dBm

	
	
	10 MHz to 3.25 GHz
	DANL to 30 dBm

	Attenuator
	Input Attenuator Range
	0 to 50 dB, in 1 dB step
	Auto or manual setup

	Maximum Safe Input Level
	Average Total Power
	≤ +33 dBm
	Input attenuator 
≥10 dB

	
	DC Voltage
	± 50 V
	

	1 dB Gain Compression
	Total Power at 1st Mixer
	> 0 dBm
	Typical;Fc  ≥ 50 MHz; preamp. off

	
	Total Power at the Preamp
	> -22 dBm
	Typical;Fc  ≥ 50 MHz; preamp. on

	
	
	mixer power level (dBm)= input power (dBm)-attenuation (dB)

	Displayed Average Noise Level (DANL)
	Preamp off
	0 dB attenuation; RF Input is terminated with a 50Ω load. RBW 10 Hz; VBW 10 Hz; span 500 Hz; reference level =  -60dBm; trace average ≥ 40

	
	9 kHz to 100 kHz
	< -93 dBm
	Nominal

	
	100 kHz to 1 MHz
	< -90 dBm - 3 x (f/100 kHz) dB
	

	
	1 MHz to 2.7 GHz
	< -122 dBm
	

	
	2.7 GHz to 3.25 GHz
	< -116 dBm
	

	
	Preamp on
	0 dB attenuation; RF Input is terminated with a 50Ω load ; RBW 10 Hz; VBW 10Hz; span 500 Hz;  reference level =  -60dBm; trace average ≥ 40

	
	100 kHz to 1 MHz
	< -108 dBm - 3 x (f/100 kHz) dB
	Nominal

	
	1 MHz to 10 MHz
	< -142 dBm
	

	
	10 MHz to 3.25 GHz
	< -142 dBm + 3 x (f/1 GHz) dB
	

	Level Display Range
	Scales
	Log, Linear
	

	
	Units
	dBm, dBmV, dBuV, V, W
	

	
	Marker Level Readout
	0.01 dB
	Log scale

	
	
	0.01 % of reference level
	Linear scale

	
	Level Display Modes
	Trace, Topographic, Spectrogram
	Single / split Windows

	
	Number of Traces
	4
	

	
	Detector
	Positive-peak, negative-peak, sample,  normal, RMS(not Video), Average(EMI), Quasi-Peak(EMI)
	Can be setup for each trace separately

	
	Trace Functions
	Clear & Write, Max/Min Hold, View, Blank, Average
	

	Absolute Amplitude Accuracy
	Absolute Point
	Center=160 MHz ; RBW 10 kHz; VBW 1 kHz; span 100 kHz; log scale;  1 dB/div; peak detector; 23°C±1°C; Signal at Reference Level

	
	Preamp off
	± 0.3 dB
	Ref level 0 dBm; 
10 dB RF attenuation

	
	Preamp on
	± 0.4 dB
	Ref level -30 dBm; 
0 dB RF attenuation

	Frequency Response
	Preamp off
	Attenuation: 10 dB; Reference: 160 MHz; 20 to 30°C

	
	100 kHz to 2.0 GHz
	± 0.5 dB
	

	
	2GHz to 3.25 GHz
	± 0.7 dB
	

	
	Preamp on
	Attenuation: 0 dB; Reference: 160 MHz; 20 to 30°C

	
	1 MHz to 2 GHz
	± 0.6 dB
	

	
	2 GHz to 3.25 GHz
	± 0.8 dB
	

	Attenuation Switching Uncertainty
	Attenuator setting
	0 to 50 dB in 1 dB step
	

	
	Uncertainty
	± 0.25 dB
	reference: 160 MHz, 10dB attenuation

	RBW Filter Switching Uncertainty
	1 Hz to 1 MHz
	± 0.25 dB
	reference : 10 kHz RBW

	Level Measurement Uncertainty
	Overall Amplitude Accuracy
	± 1.5 dB
	20 to 30°C; frequency > 1 MHz; Signal input 0 to -50 dBm; Reference level  0 to -50 dBm; 
Input attenuation 10 dB;  
RBW  1 kHz; VBW 1 kHz; after cal; Preamp Off

	
	
	± 0.5 dB
	Typical

	Spurious Response
	Second Harmonic Intercept
	
	Preamp off; signal input -30dBm; 0 dB attenuation

	
	
	+35 dBm
	Typical; 10 MHz < fc < 775 MHz

	
	
	+60 dBm
	Typical; 775 MHz ≤ fc < 1.625 GHz

	
	Third-order Intercept
	 
	Preamp off; signal input -30dBm; 0 dB attenuation

	
	
	> 1dBm
	300 MHz to 3.25 GHz

	
	Input Related Spurious
	< -60 dBc
	Input signal level -30 dBm, Att.

Mode, Att=0dB; 20-30ºC 

	
	Residual Response (inherent)
	<-90 dBm
	Input terminated; 0 dB attenuation; Preamp off

	Sweep

	Sweep Time
	Range
	204 us to 1000 s
	Span > 0 Hz

	
	
	50 us to 1000 s
	Span = 0 Hz; Min Resolution = 10 us

	
	Sweep Mode
	Continuous; Single
	

	
	Trigger Source
	Free run; Video; External
	

	
	Trigger Slope
	Positive or negative edge
	

	RF Preamplifier

	
	Frequency Range
	1 MHz to 3.25 GHz
	

	
	Gain
	18 dB
	Nominal

	
	
	
	(installed as standard)

	Front Panel Input/Output

	RF Input
	Connector Type
	N-type female
	

	
	Impedance
	50 ohm
	Nominal

	
	VSWR
	<1.6 :1
	300 kHz to 3.25 GHz; Input attenuator ≥ 10 dB

	Power for Option
	Connector Type
	SMB male
	

	
	Voltage/Current
	DC +7V / 500 mA max
	With short-circuit protection

	USB Host
	Connector Type
	A plug
	

	
	Protocol
	Version 2.0
	Supports Full/High/Low speed

	MicroSD Socket
	Protocol
	SD 1.1
	

	
	Supported Cards
	 microSD, microSDHC
	Up to 32GB capacity

	Rear Panel Input/Output

	Reference Output
	Connector Type
	BNC female
	

	
	Output Frequency
	10 MHz
	Nominal

	
	Output Amplitude
	3.3V CMOS
	

	
	Output Impedance
	50 ohm
	

	Reference Input
	Connector Type
	BNC female
	

	
	Input Reference Frequency
	10 MHz
	

	
	Input Amplitude
	-5 dBm to +10 dBm
	

	
	Frequency Lock Range
	Within ± 5 ppm of the input reference frequency
	

	Alarm Output
	Connector Type
	BNC female 
	Open-collector

	Trigger Input/ Gated Sweep Input
	Connector Type
	BNC female
	

	
	Input Amplitude
	3.3V CMOS
	

	
	Switch
	Auto selection by function
	

	LAN TCP/IP Interface
	Connector Type
	RJ-45
	

	
	Base
	10Base-T; 100Base-Tx; Auto-MDIX

	USB Device
	Connector Type
	B plug
	For remote control only; supports USB TMC

	
	Protocol
	Version 2.0
	Supports Full/High/Low

speed

	IF Output
	Connector Type
	SMA female
	

	
	Impedance
	50 ohm
	Nominal

	
	IF Frequency
	886 MHz
	Nominal

	
	Output level
	-25 dBm
	10 dB attenuation; RF input: 0 dBm @ 1 GHz

	Earphone Output
	Connector Type 
	3.5mm stereo jack, wired for mono operation

	Video Output
	Connector Type
	DVI-I ( integrated analog and digital) , Single Link. Compatible with VGA or HDMI standard  through adapter

	RS232 Interface
	Connector Type
	D-sub 9-pin female
	Tx,Rx,RTS,CTS

	GPIB Interface (Optional)
	Connector Type
	IEEE-488 bus connector
	

	AC Power Input
	Power Source
	AC 100 V to 240 V, 50 / 60 Hz

Auto range selection

	Battery Pack (Optional)
	Battery pack
	6 cells, Li-Ion rechargeable, 3S2P
	With UN38.3 Certification

	
	Voltage
	DC 10.8 V
	

	
	Capacity
	5200 mAh / 56Wh
	

	General

	
	Internal Data storage
	16 MB nominal
	

	
	Power Consumption
	<65 W
	

	
	Warm-up Time
	< 30 minutes
	

	
	Temperature Range
	+5 °C to +45 °C
	Operating

	
	
	-20 °C to + 70 °C
	Storage

	
	Weight
	4.5 kg (9.9 lb)
	Inc. all options (Basic+TG+GPIB+Battery)

	
	Dimensions
	210 x 350 x 100 (mm)
	Approximately

	
	
	8.3 x  13.8 x 3.9 (in)
	

	Tracking Generator (Optional)

	
	Frequency Range
	100 kHz to 3.25 GHz
	

	
	Output Power
	-50 dBm to 0 dBm in 0.5 dB steps

	
	Absolute Accuracy
	± 0.5 dB
	@160 MHz, -10 dBm, Source attenuation 10 dB, 20 to 30°C

	
	Output Flatness
	Referenced to 160 MHz, -10 dBm

	
	
	100 kHz to 2 GHz
	± 1.5 dB

	
	
	2 GHz to 3.25 GHz
	± 2 dB

	
	Output Level Switching Uncertainty
	± 0.8 dB
	Referenced to -10 dBm

	
	Harmonics
	< -30 dBc
	Typical, output level = -10 dBm

	
	Reverse Power
	+30 dBm max.
	

	
	Connector type
	N-type female
	

	
	Impedance
	50 ohm
	Nominal

	
	Output VSWR
	< 1.6:1
	300 kHz to 3.25 GHz, source attenuation ≥ 12 dB


